## Slide 1: Sample Preparation and Mapping
Steps
  Cleaning and drying of bulk samples
  Initial Preparation. Cutting, impregnation, cast-making, fusing
  Mounting. Stubs, embedded samples, thin sections, grain mounts
  Polishing. For quantitative analysis
  Cleaning
  Mapping
  Coating
  Handling and storage
Notes: <number>

## Slide 2: Preparing Bulk Samples
Initial Prep:
  Remove organic material with oxidizing agent such as K-permanganate or H2O2
  Rinse sediments or soils to remove soluble salts and fines if desired
  Carbonate can be removed with HCl
  Remove hydrocarbons by soaking in trichlorethane
Notes: <number>

## Slide 3
Drying:
  Wet samples must be oven- or freeze-dried before introduction into the instrument. Be aware that temperatures of over 50oC can remove structural water from clay minerals.
  Rock samples that have been cut using a saw should be thoroughly dried prior to mounting.
Notes: <number>

## Slide 4: Initial Preparation
Cutting
  Large samples must be cut to thin-section size, or cut or broken to a appropriate size for an SEM stub. Typically, diamond-blade circular saw is used. Friable samples may need to first be impregnated.
Impregnating
  Friable or porous samples should be vacuum impregnated prior to sample prep. Loose material on the sample surface can cause contamination of the electron column, and unimpregnated porous samples can cause poor vacuum by prolonged outgassing.
Cast-making
  SEM visualization of pore structure can be facilitated by vacuum impregnating a sample, followed by dissolution of the sample material using hydrofluoric acid. Topographic details of small fossils can also be facilitated by producing latex cast
Fusing
  Bulk analyses of fused rock samples can be made using EMPA techniques. Samples and standards should be fused using the same procedure.
Notes: <number>

## Slide 5: Mounting
SEM Stubs:
  SEM samples can be mounted onto Al or C stubs. C can be used for low X-ray background in the case of particulate analysis. Mounting can be done using epoxy, quick-setting glue, double-stick tape (mostly for particulates) or wax. Small grains can be mounted onto double-stick tape or partially dried carbon or silver paint.
Notes: <number>

## Slide 6: Quantitative Analysis Preparation
Embedding, Thin Section, Grain Mounts
These are the most common techniques for preparation of geological samples.
Notes: <number>

## Slide 7
Notes: <number>

## Slide 8
Sample holders
Notes: <number>

## Slide 9: Polishing
Polishing a sample to a flat, unscratched surface is CRITICAL for good quantitative analysis. An uneven, or scratched sample surface can lead to uneven production of x-rays from the sample surface, errors in absorbtion correction, and spurious results.
Notes: <number>

## Slide 10: Polishing
The polishing steps needed for a sample depends on the types of material present. Metallic phases typically polish very easily, whereas minerals with strong cleavage can be problematic. A typical polishing procedure would be as follows:
    Grinding the sample to a flat surface, using 100, 260, and 600 mesh grit diamond wheels, or equivalents.
    Wash sample in water
    Polishing the sample using 30 micron diamond grit (600 mesh) for 3 minutes.
    Ultrasonically clean sample in deionized water
    Polish the sample using 6 micron diamond grit for 3 minute
    Ultrasonically clean sample in deionized water
    Polish the sample using 1 micron diamond grit for 3 minutes
    Ultrasonically clean sample in deionized water
    Polish the sample using 0.5 micron diamond grit for 1 minute
    Ultrasonically clean sample in deionized water
Notes: <number>

## Slide 11: Cleaning
Any contaminating material, particularly skin oil, on the sample surface will end up contaminating the column. Also, contaminating material under the conductive coating of a sample can cause the conductive coating that will be placed on the sample to bubble and crack, making the sample impossible to analyze
KEEP THE SAMPLE CLEAN AND OIL FREE!!!
1. After final polish, clean the sample ultrasonically with deionized water
2. Wipe the sample surface with petroleum ether.
3. Following the petroleum ether cleaning, handle the sample as carefully as possible. Glove handling is ideal, but often inconvenient. An alternative is to handle the sample with a kimwipe, or other lint-free papr or cloth.
4. Blow the sample off with air prior to coating.
Notes: <number>

## Slide 12
Finding analysis areas on a sample using the optical microscope or electron imaging can be difficult because of the small field of view. The maximum optical field of view for our instrument is 1750 microns (1.75 mm), and the maximum electron image field is around 2500 microns (2 mm). So, having a good sample map, particularly for quantitative analysis, is very important. Sample mapping is much less important for SEM work.
Two types of map may be useful, depending on you sample.
1. Macro-map. Map of entire section. Depending on sample type, this may be hand-sketched, made with a slide copier, or with an enlarging xerox machine.
2. Micro-map. Map of areas of analytical interest. This needs to be produced with a camera-equipped microscope, and can be used to locate and document exact analytical spots. A micromap can also be produced using the microprobe.
You can also mark areas of interest on your sample, but this must be done on a clean sample surface prior to applying the conductive coat.
Mapping
Notes: <number>

## Slide 13: Sample Coating
Many geological samples are nonconductors of electricity. Therefore, if an uncoated sample is placed in the path of the electron beam, the sample will charge, causing deviation of the electron beam, as well as catastrophic decharging. The sample must be coated with conductive material, such as carbon, gold, or gold-palladium alloy.
CARBON:
For quantitative analysis and X-ray mapping, carbon is the coat of choice. Because of its low Z, it has a minimal effect on the X-ray spectrum, either in terms of producing X-ray lines or absorbing X-rays. A carbon coat is applied using a vacuum evaporator at pressures of less than 10-4 torr. The ideal coating thickness is 20 nm.
GOLD:
Gold can be a better choice for SEM coating because of its higher secondary electron yield.
Notes: <number>