## Slide 1: Factors affecting CT image RAD 323 2014
Alhanouf Alshedi
Email: aalshedi@ksu.edu.sa

## Slide 2: CT Image Quality
CT image is influenced by several parameters, of which some depend on the performance of the CT scanner and some depend on the operator`s selection of parameters.
Parameters that depend on CT machine: pixel size, temporal resolution and reconstruction algorithms.
Parameter that depend on operator : kV, mAs and slice thickness.

## Slide 3: Focal spot size
CT utilizes different focal spots sizes.
Focal spot size Is determined by:
Filament size
Filament length
SMALLER FOCAL SPOT - Low mA
SMALLER FOCAL SPOT – sharper image

## Slide 4: mA – tube current
mA is the number of electrons emitted or flowing from cathode to anode.
S is the time of exposure (seconds). - mAs is the tube current(milli-amper) for a certain length of time (second).

## Slide 5: mA
CURRENT
INTENSITY
ENERGY – NO CHANGE
2 * mA = 2 * number of photons
4 * mA = 4 * number of photons

## Slide 6: Cont.
To improve image we need to reduce motion & noise
Avoiding motion – mA time
Pediatric technique modification
Reducing noise - mAs
MOTION

## Slide 7: kVp
Potential difference between cathode and anode (kiloVolts).
kVp
ENERGY
INTENSITY
15% INCREASE OF KVP = 2 * mAs

## Slide 8: kVp in CT
Usually in the range of 80-140 kV
Too low KV noise
(not enough penetration of the patient )
Too high kV over exposure

## Slide 9: Filtration
Filter
Patient
Detector

## Slide 10: Cont.
Filtration removes long-wavelength x-rays that do not play a role in image formation but cause pt dose. Energy of beam and beam becomes “harder”.
Filtration uniforms the energy of beam.
CT filters are usually 3mm, added flat or shaped copper filters can range from 0.1 -0.4 mm.
Special filters such as “bowtie” made of Teflon can reduce beam hardening artifacts.

## Slide 11: Filtration effect
FILTRATION
INTENSITY
ENERGY

## Slide 12: Collimation
Filter
SHAPES BEAM +
REDUCES AMOUNT OF
SCATTER RADIATION
Patient
DEFINES SLICE
THICKNESS +
REDUCES SCATTER
RECHING THE PATIENT
Detector

## Slide 13: Collimator
Collimation removes scatter radiation improves resolution.
Some scanners include an anti-scatter grid placed in front of detectors to remove scatter radiation and improve image quality.

## Slide 14
Steps of CT image formation
CT image formation
Data acquisition
Image reconstruction and processing
Image display and storage

## Slide 15: Digital & Analog images
Analog images: are continuous images e.g black & white chest x-ray, because they represent continuous distribution of light intensity as a function of position.
Digital images: are numerical representation of objects. The formation of digital images requires a digital computer.
Any information that enters the computer for processing must converted into digital form or numbers.

## Slide 16
Analog image
A/ D converter
Image reconstruction and processing
D/A converter
Analog image

## Slide 17
Image domains
Image domains:
images can be represented
into 2 domains:
spatial domain
Frequency domain
Digital imaging can transform an image from the spatial domain into frequency domain using a Fourier transform (FT). The FT ˉ¹ (inverse FT) is used to transform the image from frequency domain back into spatial domain.
Spatial domain
Frequency domain

## Slide 18
Radiography and CT acquire images in the spatial domain.
MRI acquires images in the frequency domain.
The major reason for using frequency domains is to enhance or suppress certain features of the image.
Some image processing operators are more efficient or only practical when applied in the frequency domain.
High spatial frequencies (image detail)
Low spatial frequencies (image contrast)

## Slide 19
Digital image
A digital image is a representation of a two-dimensional image as a finite set of digital values, called picture elements or pixels

## Slide 20
An image is represented by a number of picture elements (pixels). These pixels are arranged into rows and columns. Y representing the columns and X representing the rows.
0 1 2 3 4 5 6 7 8 9 10 11
0
1
2
3
4
5
6
7
8
8, 3

## Slide 21: Matrix and FOV
The matrix is a digital image made up of two dimensional arrays. It consists of columns (M) and rows (N). The matrix size is related to the FOV. It can be determined by :
Matrix size= M x N x K bits
Where k bit is the bit depth (each pixel will have 2 gray levels).
If M=N the image is square. If M ≠ N then the image is rectangle.

## Slide 22
Each pixel contains a number that represents the brightness level (gray level). This number represents tissue characteristics, in x-ray and CT, whereas, in MRI it represents proton density and relaxation times.
Pixel size can be calculated using:
Pixel size = FOV/ matrix size
The larger the matrix size the smaller the pixel size the better the resolution if FOV is constant.
Voxel is the representation of volume (thickness)

## Slide 25: Why do we need to digitise images
Images need to be digitised to perform several fundamental operations:
Image enhancement
Image restoration
Image analysis
Image compression
Image synthesis

## Slide 26
1- Image enhancement:
To produce an image the is more pleasing to the eye. Shapes and edges can be enhanced to improve quality of image. This operation includes: contrast enhancement, edge enhancement, spatial and frequency filtering, noise reduction.
Noise reduction

## Slide 27
2- Image restoration:
To improve quality of distorted, degraded or blurred (from motion) images by compensating or undoing the defect using special filters.

## Slide 28
3- Image analysis:
Allows measurements and statistics to be performed in addition to image segmentation, feature extraction and classification of objects.
4- Image compression:
- For large amounts of data, compression is needed to reduce size and facilitate processing, transmission and storage.
- Compression can be 2 types:
A- Lossy compression some loss of detail when image is decompressed, provides higher levels of data reduction.
B- Lossless compression no loss of information when image is decompressed, used for medical imaging.

## Slide 29: Image compression
5- Image synthesis:
Create images from other images or non-image data. Ex. Reconstruction that are the base of CT and MRI and 3D techniques.

## Slide 30: Any Question?
Thank You