## Slide 1: CT Physics and Instrumentation RAD 323 2014
Alhanouf Alshedi
Email: aalshedi@ksu.edu.sa

## Slide 2
Mid terms 1 and 2 | 20+20=40
Report/assignment | 10
Case presentation | 5
Attendance | 5
Final | 40
Total | 100

## Slide 3: Topics to be covered
1st Lecture: Introduction to C.T and Historical prospective, CT physical principle.
2nd Lecture: CT instrumentation and x-ray system and Data acquisition .
3rd Lecture: Detectors and detector characteristics.
4th Lecture: Factors affecting CT image, Digital fundamentals and Image domains. 1st Midterm Exam
5th Lecture: Image reconstruction.
6th Lecture: Image Post-processing and Manipulation tools and Windowing.
7th Lecture: CT Computer and Image processing system.
8th Lecture: Quality control and Patient dose. 2nd Midterm Exam.
9thLecture: Advances in CT slip ring technology.
CT visit will be scheduled with KKUH.
Student`s Presentations .

## Slide 4: Useful Books
E. Seeram, Computed Tomography: Physical Principles, Clinical Applications, and Quality.
W. A. Kalender, Computed Tomography: Fundamentals, System Technology, Image Quality.

## Slide 5: Introduction to Computed Tomography and Historical Prospective
1st Lecture
Alhanouf Alshedi
Email: aalshedi@ksu.edu.sa

## Slide 6: What is Tomography?
Tomography comes from the Greek word tomos meaning “section” or “cut” and graph meaning write. A conventional tomogram is an image of a section of a patient parallel to the film.
With the development of technologies, parallel sections were overcome by cross-sections and conventional tomography was replaced with computerised tomography C.T.

## Slide 8: How are x-rays produced?

## Slide 10
Production of X-rays
The filament is heated by passing an electrical current through it. The filament then emits electrons by thermionic emission.
The electrons are then repelled by the negative cathode and attracted by the positive anode and bombard the target.
Each electron arrives at the target with a kinetic energy, as the electrons penetrate several micrometers into the target, they lose their energy by several processes.

## Slide 11: Cont.
Heat.
Bremsstrahlung radiation.
Characteristic x-rays.
Total number of electrons converted to heat is 99% and only 1% of the electrons are converted to x-rays
Heat:
When an electron interacts with an outer loose electron, it loses a small amount of its energy. This results in unwanted heat.

## Slide 12: 2. Characteristic X-rays
The high energy electron can also cause an electron from an inner shell to be knocked out from its place.
This vacancy is filled by an electron further out from the nucleus. The well defined difference in binding energy, characteristic of the material, is emitted as a monoenergetic photon. When detected this X-ray photon gives rise to a characteristic X-ray line in the energy spectrum.

## Slide 14: 3. BREMSSTRAHLUNG RADIATION
An incoming free electron gets close to the nucleus, the strong electric field of the nucleus will attract the electron, changing direction and speed of the electron. The Electron looses energy which will be emitted as an X-ray photon.
X-rays originating from this process are called bremsstrahlung. Bemsstrahlung is a German word meaning: "Bremse" means "brake and "Strahlung" means "radiation“.

## Slide 16: What is CT?
In principle, Computed Tomography (CT) measures the attenuation of beams passing through sections of the body from hundreds of different angles, and then from these measurements a computer is employed to reconstruct pictures of the bodies interior.
CT uses the same principle as radiography, i.e x-ray beam passes through the pt body.
CT produces a volume of data that can be manipulated in order to demonstrate various bodily structures based on their ability to attenuate the X-ray beam.

## Slide 18: Cont.
While a typical digital image is composed of pixels (picture elements), a CT slice image is composed of voxels (volume elements). Taking the analogy one step further, just as a loaf of bread can be reconstituted by stacking all of its slices, a complete volumetric representation of an object is obtained by acquiring a contiguous set of CT slices.

## Slide 19: CT Image
Brain CT images.

## Slide 20: Cont.
The gray levels in a CT slice image correspond to X-ray attenuation, which reflects the proportion of X-rays scattered or absorbed as they pass through each voxel.
X-ray attenuation is primarily a function of X-ray energy and the density and composition of the material being imaged.

## Slide 21: How it works?

## Slide 22: Why do we need CT?
Limitations of radiography:
Superimposition of structures, especially structures that only differ slightly in density e.g. tumor (limited contrast resolution).
Qualitative rather than quantitative.
Limitations of conventional tomography:
Contrast degradation.
Image blurring.
Limitations of film itself as a detector.

## Slide 23
How does C.T overcome the limitations of radiography and tomography?
X-ray beam is transmitted through a specific cross-section removes superimposition.
X-ray beam is tightly collimated scatter and contrast.
X-ray beam is detected by special detectors which are quantitative & can measure subtle differences in tissue contrast.
Ability to manipulate and adjust image after scanning (digital technology)

## Slide 24: Any Question?
Thank You