## Slide 1: Instrumental Methods: Intro
 Types of Instrumental Methods
 Fundamental Components of an Instrument
 Instruments Measure Voltages and Currents!
 Basics of Analytical Methods
Review
Terminology
Some notes and figures in this course have been taken from Skoog, Holler and Neiman, Principles of Instrumental Analysis, 5th Edition, Saunders College Publishing.

## Slide 3: Basic Instrument Components
Source: produces some form of energy or mass that is relevant to the measurement at hand
Sample Holder or “Cell”: contains the sample with your analyte of interest
Discriminator: selects the desired signal from the source or the sample
Input Transducer: detects the signal from the sample, source or discriminator. AKA the detector.
Processor: manipulates the signal electronically or mechanically to produce some useful value
Readout: displays the signal in some useful form.

## Slide 5: Instruments Measure 1 of 2 Things.
VOLTAGE (V), volts, electrical potential across two electrodes.
Current (A), amperes, the flow of electrons across some point.
V = IR
R= resistance in Ohms
THEY DO NOT MEASURE ABSORBANCE, FREQUENCY, MASS, CONCENTRATION OR ANY OTHER CONVENIENT UNIT. THEY MEASURE CURRENTS OR VOLTAGES, WHICH ARE THEN CONVERTED TO THESE CONVENIENCE UNITS!

## Slide 6
Basic Questions Regarding All
Analytical & Instrumental Methods
Defining the “instrumental analysis” Problem:
o What accuracy and precision are required?
o How much sample do I have available, and how much money do we have available for the analysis?
o What concentration is the analyte present at and can we pre-concentrate or dilute the sample?
o What interferences might be present and can we eliminate or mask them?
o What are the properties of the sample matrix?
How much time do I have (time = money)?

## Slide 7
Some Basic Definitions (Review)
A sample is collected or taken
An aliquot is usually selected from the larger, bulk sample for preservation, preparation and/or analysis
A technique implies the use of a specific type of instrument for analysis
A method is the procedure followed when utilizing an instrumental technique
A protocol is a regulatory or “officially” recognized method that must be adhered to
GLP stands for Good Laboratory Practice
GMP stands for Good Manufacturing Practice

## Slide 8: Relevant Analytical Parameters
These are new. You should be familiar with accuracy, precision, average, standard deviation, %relative standard deviation, etc.
Analytical Sensitivity: The slope of the calibration curve (IUPAC Definition)
Thus, other factors being equal, the method with the steepest calibration curve will be more “sensitive”
  Better ability to discriminate between numerically close concentrations.

## Slide 10: Detection Limit (DL, LOD, MDL)
Most widely disputed term in instrumental methods.
The minimum concentration of analyte that can be detected, based on the analytical signal.
  DETECTED, not necessarily known with any great confidence!
  LOD = [Mean Blank Signal] + 3 x Std. Dev. Blank Signal
In general, 3 is chosen as the multiplier because at + 3 STDEV, you are 99% confident you are not measuring signal from noise, background, etc.
Measurements at or near the limit of detection are not necessarily precise (high %RSD)! This is what instrument manufacturers will quote you, as measured under the most ideal, not regularly attainable, conditions!
The STDEVBlank signal is often replaced with the standard deviation for some very, very low (near the DL) sample you have prepared.
This signal is then used with the cal. curve to calculate a DL.

## Slide 11: Calculate the limit of detection for the determination of yttrium by the PINK METHOD if the measurement of the blank 50 times yielded a mean signal of 1.00 and a standard deviation of 0.20 absorbance units.

## Slide 12: Limit of Quantitation (LOQ)
Another somewhat disputed term.
The LOQ is generally considered the minimum concentration of analyte that can be “accurately” and “precisely” determined. Exact definitions vary, however..
You measure a blank AND a VERY low concentration sample that is near the detection limit) numerous times, and then use that data.
  10 times is the typical number of replicates
This signal is used in the calibration curve to calculate the MDL.

## Slide 13: Dynamic Range
Usually called the Linear Dynamic Range, this is the concentration range over which the calibration curve has a linear shape.
You have probably seen an instrument exceed its linear dynamic range with the SPEC 20
  Beer’s Law fails at increasing concentrations due to instrumental and sample solution factors.
Sample matrix, analyte and method dependent.
You usually want to work with linear calibration curves if at all possible (much less complex than quadratic, exponential or polynomial fits)
  Determination of Metals by AAS : 1-3 orders of magnitude
  Determination of Metals by ICP-AES: 5-8 orders of magnitude

## Slide 15: Selectivity
Also known as discrimination
The ability to discern different, yet closely spaced analytical signals.
The spectrometer on the SPEC 20 can discriminate wavelengths of light that are about 20 nm apart (even if you can set wavelengths only 5 nm different)
The spectrometer on our Varian ICP can discriminate wavelengths of light that are 0.005 nm apart!
Better selectivity means you can be sure which signal is which when you have more than one analyte in the sample!
  However, if all other conditions are equal, increasing selectivity will decrease the amount of signal you can measure (reduce the LOD)!

## Slide 16: Bandwidth is closely related to selectivity in optical spectrometers. It is a measure of what range of light we allow to strike the detector at any given time.

## Slide 17: EVERYTHING YOU DO IN THIS CLASS WILL BE A BATTLE! THE BATTLE BETWEEN SIGNAL AND SELECTIVITY! There is no way to maximize both. You have to choose some happy medium, where you get enough signal to detect the analyte, but can also be selective enough so that you are sure of what you are detecting.