## Slide 1: Nephelometry and Turbidimetry
Cls 332 lab 4

## Slide 2: introduction
Turbid metric + nephlometeric analysis:
When part of incident radiant energy is dissipated by absorption, reflection, and refraction, while the remainder is transmit light as a function of the concentration of the dispersed phase is the basis of turbid metric analysis

## Slide 4
 | Nephlometer | Turbidimeter
Definition | the measurement of the intensity of scattered light at right angles to the direction of the incident light as a function of the concentration of the dispersed phase ,It is most sensitive for very dilute suspensions (100 mg/L). | Light passing through a medium with dispersed particles, so the intensity of light transmitted is measured.
Instrument used | Nephlometery machine | spectrophotometer
Type of light measured | Scattered light | Transmitted light
Arrangement of photometer | measure the light scattered at right angle to the direction of the propagation of light from the source. It could be movable detectors which allow operator to vary the angle of detection | made in the same direction as the propagation of the light from the source.
Clinical uses | Ag-Ab rxn, immunocomplex rxn,ppts, lipoprotein | Ag-Ab rxn, immunocomplex rxn,ppts, liver dis, protein in urine or CSF

## Slide 5
Nephelometry and Turbidimetry

## Slide 6: Instrumentation:
1- light source:
Tungsten its relatively low intensity makes it less useful for samples with low light scattering. Alternatives are:
quartz halogen lamp, xenon lamp and laser which have higher intensities than tungsten lamp.
2-lens assembly:
Light enter the sample holder through lens assembly.

## Slide 7
3-there is provision for the insertion of filter between the sample and source of light(monochromator).
4- detector (photo –cell):
It is shielded to minimize interference from stray light.
5-Read out device:
Light intensity is converted to an electrical signal by the detector .

## Slide 8: Example:
The turbidity of a dilute barium sulphate suspension.
The concentration of the reactants must be controlled by adding pure solid barium chloride of definite grain size.
-NACL and HCL are added before the precipitation in order to inhibit the growth of microcrystal of barium sulphate
-A glycerol ethanol solution helps to stabilise the turbidity.
The reaction vessels is shaken gently in order to obtain a uniform particle size.
Each vessel should be shaken at the same rate and the same number of times.

## Slide 10: MEASURMENTS:
1-plug the instrument into ground outlet.
2- choose desirable scale from 0-10
starting with the highest conc. (for std 1=scale 10)
3- turn power switch on.
4-selecte desirable range by range selector at desirable position .
5- select filter required.

## Slide 11
6- transfer your standards in the cleaned cell and place them in cell holder.
7- remove the standards.
8- fill the second cell with blank to set zero .
9- check the reading of the standards again.
10- measure your unkown.
11- draw calibration curve.

## Slide 12: Precaution:
-Number and size of the particles should remain constant if repeated preparation are made
-clean cell & filter
- avoid air bubbles (high reading).
-dilute sample if there is need.
-prepare the blank ,standards, Sample at the same time.(to avoid ppt)

## Slide 13: Advantages and disadvantages :
Advantages : rapidity of procedure and simplicity of the measurements.
Disadvantages: lack of accuracy of the measurements.