## Slide 1: Nephelometer and Turbidimeter

## Slide 2
Introduction

## Slide 3
Turbidimetry
Method for determining the amount of cloudiness, or turbidity, in a solution based upon measurement of the effect of this turbidity upon the transmission and scattering of light.

## Slide 5
Turbidity in a liquid is caused by the presence of finely divided suspended particles..
Turbidimetry

## Slide 6
If a beam of light is passed through a turbid sample, its intensity is reduced by scattering, and the quantity of light scattered is dependent upon the concentration and size distribution of the particles.
Turbidimetry

## Slide 7
Is the measurement of the intensity of light transmitted through the sample, the unscattered light, is measured.
.
Turbidimetry

## Slide 8
Turbidimetry is the technology used to protein assays such as Apolipoproteins, Lp(a), CRP, RF, ASO, C3, C4, Immunoglobulins etc...
Turbidimetry

## Slide 9
Nephelometry

## Slide 10
Nephelometry
Is defined as the detection of light energy scattered towards a detector that is not in the direct path of the transmitted light.

## Slide 11
Common nephelometers measure the intensity of the scattered light at right angles to the incident light, but some have detectors placed at angles of 60-70.
Nephelometry

## Slide 12
This enables to take advantage of the increased forward scatter intensity caused by light scattering from large particles.
e.g: The amount of light scattered is proportional to the concentration of antigen or antibody(protein) in the solution
Nephelometry

## Slide 13
‘Tyndall effect‘
Means the phenomenon in which light is scattered by very small particles in its path; (or in colloid systems, such as suspensions or emulsions) it makes a beam of light visible; the scattered light is mainly blue.
It is named after the 19th century physicist John Tyndall.

## Slide 14
also known as Tyndall scattering

## Slide 15
When Tyndall effect happen?
When the suspension is viewed at right angle to the direction of the incident light the system appears shining due to the reflection of the light from the particles
of the suspension

## Slide 16
Turbidity can be measured on most routine analysers by a spectrophotometer (absorbed light)
The intensity of scattered light is normally measured by nephelometer

## Slide 17
The main uses of nephelometers relate to air quality measurement for pollution
monitoring.
Biological contaminants include mold, fungus, bacteria, viruses..
Nephelometry

## Slide 19
Instrumentation:
1- light source:
Tungsten its relatively low intensity makes it less useful for samples with low light scattering. Alternatives are:
quartz halogen lamp, xenon lamp and laser which have higher intensities than tungsten lamp.

## Slide 20
2-lens assembly:
Light enter the sample holder through lens assembly.

## Slide 21
3- monochromator
Is provision for the insertion of filter between the sample and source of light.

## Slide 22
4- detector (photo –cell):
It is shielded to minimize interference from stray light to measure light reflected (nephelometer)
It could be movable detectors
which allow operator to vary the angle of detection.

## Slide 23
Read out device:
Nephlometer measure the intensity of a scattered light. Light intensity is converted to an electrical signal by the detector .

## Slide 24
Comparison of Nephelometry and Turbidimetry :
Nephelometry methods are more sensitive than turbidimetric methods, with a detection limit of approx. 10ug/mL whereas the detection limit of Turbidimetry methods are 20 to 30 ug/mL.

## Slide 25
Nephelometry shows a slightly better detection limit as compared to turbidimetry in case of lipemic samples and also in case of pure media such as CSF

## Slide 26
Nephlometry is differs from turbidimetery in the arrangement of the photometer, since turbidity measurements are made in the same direction as the propagation of the light from the source.
While the nephlometry is measure the light scattered at right angle to the direction of the propagation of light from the source

## Slide 27: Criteria for analysis
1- Number and size of the particles should remain constant if repeated preparation are made.
2- The precipitate must be fine enough so as not to settle rapidly.

## Slide 28
Advantages and disadvantages :
Advantages are : rapidity of procedure and simplicity of the measurements.
Disadvantages: lack of accuracy of the measurements.

## Slide 29: Principle
The turbidity of a dilute barium sulphate suspension is difficult to reproduce so we have to follow procedure accurately.
The concentration of the reactants must be controlled by adding pure solid barium chloride of definite grain size.

## Slide 30
Sodium chloride and hydrochloride acid are added before the precipitation in order to inhibit the growth of microcrystal of barium sulphate
A glycerol ethanol solution helps to stabilise the turbidity.
The reaction vessels is shaken gently in order to obtain a uniform particle size. Each vessel should be shaken at the same rate and the same number of times.

## Slide 31: What do we need??
1- Preparation of standards = 1 mg/ml.
1.814 g of potassium sulphate dissolve in 1 litter of H₂O. Or
0.3 g of p. Sulphate to 300 ml of H₂O.
2-NACL-HCL:
240 g of NACL + 20 ml of conc. HCL to 1 litter of water.
3-Glycerol-ethanol
100 ml of glycerol +200 ml of absolute alcohol
4-barium chloride
Dissolve in the weighing boat before addition of water which is (100 ml)

## Slide 32: Procedure
1- Preparation of standards in (ml)
 | blank | Std 1 | Std 2 | Std 3 | Std 4 | Std 5 | Std 6
Conc. of potassium sulphate std(1 mg/ml)=0.001 g/ml | ...... | 0.5 | 1 | 1.5 | 2 | 2.5 | 3
NACL-HCL | 10 | 10 | 10 | 10 | 10 | 10 | 10
GLYCEROL-ETHANOL | 20 | 20 | 20 | 20 | 20 | 20 | 20
BARIUM CHLORIDE | 0.3 g | 0.3 g | 0.3 g | 0.3 g | 0.3 g | 0.3 g | 0.3 g

## Slide 33
MEASURMENTS:
1- Plug the instrument into ground outlet.
2- Choose desirable scale starting with the highest conc. (scale 10 is good).
3- Turn power switch on.
4- Select desirable range by range selector at desirable position .
5- Select filter required.

## Slide 34
6- Transfer your standards in the cleaned cell and place them in cell holder.
7- Adjust the standards control to obtain the maximum reading.
8- Remove the standards.
9- Fill the second cell with blank to set zero .
10- Check the reading of the standards again.
11- Replace the std. With other std (less conc.) and note the various scale reading.
12- Draw calibration curve.

## Slide 35
Precaution:
1- Clean cell.
2- Clean filter
3- Avoid air bubbles (high reading).
4- Dilute sample if there is need.
5- Prepare the blank ,standards, sample at the same time.

## Slide 36
Clinical uses:
1- Nephlometery:
Study of lipoprotein, determination of specific protein by immunological methods.
2- Turbidimetry:
Liver disease and protein contents of urine and CSF.