## Slide 1: Radionuclide Generators
Lab 8

## Slide 2: Generators
Properties:
A radionuclide generator must be sterile and pyrogen-free.
  The generator system may be sterilized either by
    autoclaving the entire column
    Or by preparing it from sterile materials under aseptic conditions.
  To maintain sterility
    Bacteriostatic agents are added to the generator column.
    Or a membrane filter unit is attached to the end of the column.

## Slide 3: An ideal radionuclide generator should be:
Simple
Convenient
Rapid to use
Give a high yield of the daughter nuclide repeatedly and reproducibly
Properly shielded to minimize radiation exposure
Eluate should be free from the parent radionuclide and the adsorbent material
Other extraneous radioactive contaminants should be absent in the eluate.Convenient
The daughter nuclide should decay to a stable or very long-lived nuclide so that the radiation dose to the patient is minimal

## Slide 4: Important Radionuclide Generators 99Mo–99mTc Generator
99Mo radionuclide B‾emission 87% metastable state 99mTc
13% ground state 99Tc
99mTc has a half-life of 6 hr and decays to 99Tc by isomeric transition
Ground state 99Tc has a half-life of 2.1X105 years and decays to stable 99Ru by B‾ emission

## Slide 5: Construction
Liquid Column Generator
Solid Column Generator

## Slide 6: Construction
Liquid Column Generator
20% NaOH solution of 99Mo in a glass column
99mTc radioactivity extraction with
methyl ethyl ketone (MEK)
Organic phase will be evaporated
and the 99mTcO4‾ dissolved in
isotonic saline for clinical use.

## Slide 7: Construction
Liquid Column Generator
Advantage
  The cost of 99mTc is low.
Disadvantage
  Needs a lot of manipulation in the overall method.
It is rarely used in nuclear medicine.

## Slide 8: Construction
Solid Column Generator
The 99Mo–99mTc or ‘‘Moly’’ generator is constructed with alumina (Al2O3) loaded in a plastic or glass column.
The 99Mo radioactivity is adsorbed on alumina in the chemical form MoO24‾ (molybdate).
The 99mTc radionuclide is eluted as sodium pertechnetate (Na99mTcO4) with a 0.9% NaCl solution.

## Slide 9: Construction
Solid Column Generator
Shielding
The generator columns are shielded with lead for radiation protection.
Some commercial firms use depleted uranium in lieu of lead for shielding high 99Mo activity because 238U has higher Z and therefore attenuates g rays more efficiently
Depleted uranium is natural uranium from which 235U has been removed, leaving only 238U).

## Slide 10: Construction
Solid Column Generator
Yield of 99mTc
One can calculate the theoretical yield of 99mTc from a Moly generator at a given time by using this Eq.
For practical reasons, it is not possible to obtain a complete yield of 99mTc from a generator as predicted by this Eq.

## Slide 11: Construction
Solid Column Generator
Yield of 99mTc
One can calculate the theoretical yield of 99mTc from a Moly generator at a given time by using this Eq.
For practical reasons, it is not possible to obtain a complete yield of 99mTc from a generator as predicted by this Eq.

## Slide 12: Construction
Solid Column Generator
Yield of 99mTc
For practical reasons, it is not possible to obtain a complete yield of 99mTc from a generator as predicted by this Eq.
  Why?
The yield may be reduced by
Column defect such as
channeling in the adsorbent bed
or by autoradiolysis due to high radioactivity whereby the chemical form of 99mTc changes.

## Slide 13: Construction
Solid Column Generator
Yield of 99mTc
Problem
A 2.6-Ci (96.2-GBq) Moly generator calibrated for Wednesday noon was received on Tuesday before. What would be the total 99mTc activity eluted at 8:00 a.m. on Friday?

## Slide 14: Quality Control of 99mTc-Eluate
Since 99mTc activity is used for humans, several quality control tests of the 99mTc-eluate are mandatory.
  99Mo Breakthrough
  Aluminum Breakthrough
  pH
  Radiochemical Purity

## Slide 15: 99Mo Breakthrough
Originates from the small quantity of 99Mo that may be eluted with 99mTc.
The US Pharmacopeia limit is 0.15 mCi 99Mo/mCi 99mTc per administered dosage at the time of administration.
The 99Mo contamination is measured by detecting 740-keV and 780-keV photons of 99Mo in a dose calibrator or a NaI(Tl) detector coupled to a pulse height analyzer.

## Slide 16: Aluminum Breakthrough
Originates from the alumina bed of the generator.
The presence of aluminum in the 99mTc-eluate interferes with the preparation of 99mTc-sulfur colloid; particularly phosphate buffer in colloid preparations tends to precipitate with excessive aluminum.
The USP limit is 10 µg Al/ml 99mTc for fission-produced 99Mo.
Detected by the colorimetric method using aurin tricarboxylic acid or methyl orange, and can be quantitated by comparison with a standard solution of aluminum.
Excessive amounts of aluminum in the eluate indicate lack of stability of the column.

## Slide 17: pH
The pH of the eluate should be between 4.5 and 7.5.
Can be checked quantitatively with a pH meter or qualitatively with pH paper.

## Slide 18: Radiochemical Purity
The impurities of the 99mTc eluate are different chemical forms of radioactivity other than 99mTcO-4
These impurities should be checked by suitable analytical methods.

## Slide 19
Thank You