Lesson 4: Hearing Aid Components

Block Diagrams:

A block diagram shows what operations a device carries out on signals within the hearing aid, and in what order each of these operations is carried out. This helps the clinician understand what effect varying a control will have, and just as importantly, what effects it will not have.

Microphones:

The function of the microphone is to convert sound into electricity. Because it changes energy from one form to another it is known as a transducer. For a perfect microphone, the waveform of the electrical signal coming out of the microphone is identical to the waveform of the acoustical signal going into the microphone.

Microphones act in a linear fashion, so every time the pressure of the input signal doubles, for instance, the output output voltage also doubles. The relationship between the size of the output voltage and the size of the input sound pressure is known as the sensitivity of the microphone.

The microphone is basically a closed box that is divided into two small volumes by a thin polymer diaphragm. Sound pressure enters the microphone through a small tube shown at the left, and then travels to the region called the “front volume” of the microphone. In the front volume, the sound pressure creates a small motion of the diaphragm. On the other side of the diaphragm, the “back volume” contains a metal plate that is coated with an electret material that holds a permanent electrical charge. The motion of the diaphragm near the charged back plate creates a small electrical signal that is amplified to become the electrical output signal of the microphone.

Microphone Imperfections:

The major imperfection with microphones is that they eventually break down if they are exposed to adverse chemical agents, like perspiration. Also all electronic components generate small amounts of random electrical noise, and microphones are no exception. The noise is partly the result of random motion of air molecules against the diaphragm, and partly the result of random electrical activity within the internal microphone amplifier. The last imperfection is that microphones are subject to wind noise. When wind hits an obstacle like a head, a pinna, or a hearing aid, turbulence is created.

Directional Microphones:

Directional microphones have been available for use in hearing aids since the early 1970s. The clinical benefit of improved hearing in noise with a directional microphone has been understood since at least the 1980s. Nonetheless, hearing aids with directional microphone responses did not gain significant market acceptance until the mid-1990s. Why is that?

Part of the reason, certainly, is improvement in the design and directional performance in the more recent systems. However, directionality also has disadvantages in some listening situations. Modern directional systems all provide a method of switching between a directional mode and a non-directional mode so that the wearer can easily put the hearing aid in the appropriate mode for each listening situation. Many researchers believe that the primary reason for the wide acceptance of modern directional systems is this flexibility.

Microphones can be made to have a sensitivity that depends on direction of arrival by feeding sounds to both sides of the diaphragm from two separate inlet ports. Directional microphones suppress noise coming from some directions, while retaining good sensitivity to sounds arriving from one direction.

The directional sensitivity of a microphone is usually indicated on a polar diagram. A whole family of response shapes can be generated using directional microphones. As the shape moves from cardiod to a supercardiod, to a hypercardiod, the sensitivity to sounds from the back grows, but the sensitivity to sounds from the sides diminishes. In the extreme case, which is referred to as a figure-8 or bi-directional pattern, the front and rear have the same sensitivity, but the microphone is completely insensitive to sounds coming from the side. The other extreme is an omni-directional (non-directional) microphone, which has a single port and a polar diagram in the shape of a circle.

What response shape is most desirable? In many real life situations, unwanted noise arrives more or less equally from all directions, because even if the noise originates from one or two sources, room reflections cause the energy to arrive at the aid wearer from all directions. By contrast, if the aid wearer is standing close to the person he or she wants to hear, the wanted signal will arrive mostly from directly in front.

A good directional microphone should therefore have maximum sensitivity for sounds arriving from directly ahead, but the sensitivity averaged across all other possible directions should be as low as possible, if intelligibility in noisy environments is to be maximized.

The ratio of sensitivity for frontal sounds relative to sensitivity averaged across all other directions is referred to as the directivity index (DI). It is expressed in dB.

Sometimes the ratio of frontal sensitivity to rearward sensitivity is quoted in hearing aid specifications. This front-to-back ratio is a misleading measure because it says nothing about the effectiveness of the hearing aid in suppressing noise arriving from directions other than precisely behind the aid wearer. For the different shapes, the highest three dimensional directivity index equal to 5.9 dB is actually obtained with the hyper-cardioid which has a poor front-to-back ratio

The directional microphone is very similar to the omnidirectional microphone described above except that it has a second sound entry port. In use, the two ports are normally aligned horizontally, along a line that points in the direction the user is facing. Thus, one port is called the “front port” and the other the “rear port” or “back port.” As figure 2 shows, the front port brings sound into the front volume of the microphone and the back port brings sound into the back volume. This rear port often contains a screen resistance, whose purpose will be discussed later. Since both sides of the diaphragm now have an acoustic pressure, the diaphragm motion is driven by the difference in acoustic pressure between the front and back volumes. Consequently, a directional microphone of this type is often called a pressure-difference microphone or a pressure-gradient microphone.

Figure 2. In a simple microphone, sound pressure enters the front volume and causes a small vibration

of the diaphragm. The directional microphone has two ports. The diaphragm is driven by the pressure

difference in the front and back volumes.

If the acoustic pressure in the front and back volumes were the same, then the pressure difference would be zero and the microphone would have no output. Luckily, this does not occur because of the separation between the two sound entry ports. This means that for sounds that originate in front of the user the sound arrives at the front port a little sooner than it arrives at the back port. This time delay causes a small phase shift between the pressure signals in the front and back volumes.

As Figure 3 illustrates, the time delay and phase shift vary for different arrival angles of sound. For each different arrival angle, there is a different time delay and a different output for the microphone. The small pattern at the center of Figure 3 shows the case where no damping screen is used in the microphone. Sounds coming directly from the side have no time delay, and therefore the microphone has no output in this direction. When a damping screen is used, its resistance value can be selected to give any of the free field patterns shown in Figure 1.

Figure 3. Sound from different directions arrives at the two ports with different time delays

and generates different microphone responses.

The subtraction process that takes place in the difference microphone

generally causes the microphone sensitivity to be significantly lower than

that of an equivalent omnidirectional microphone. Figure 4 shows the frequency response of the sensitivity of a directional microphone and of a

similar omnidirectional microphone. The decrease in sensitivity is significant, especially at low frequencies, and is greater when the separation of the two ports is smaller.

Figure 4. The frequency response for a directional microphone, or for two microphones processed to give directivity, falls at low frequency.

DUAL-MICROPHONE PROCESSING:

The other approach to achieving a directional pattern in a hearing aid

is to use two omnidirectional microphones and combine their outputs

to create the directional signal. Figure 5 shows the general form of the

processing that is needed to produce a directional pattern. The output signal of one microphone is subtracted from that of the other. The specific directional pattern may be changed among the possibilities shown in Figure 1 by adding a simple delay circuit, as shown in Figure 5.

Figure 5. A directional pattern is generated by two omnidirectional microphones whose

outputs are subtracted.

The two methods of obtaining directionality have much in common. Figure 6 shows functional block diagrams of the two methods. In the two-port directional microphone, the acoustic signals from the two ports subtract by driving the diaphragm on its opposite sides. The difference signal is then converted to an electrical signal by the microphone. In the dual microphone directional system, the acoustic signals at the two microphone ports are each converted to electrical signals and then subtracted in the electrical circuits.

Figure 6. The two ways of obtaining directionality do the same things in a slightly different

order. The two-port directional microphone does it all in the acoustics of the microphone.

The dual-microphone system does the same things electrically.

In each system, the specific pattern shape (one of the options shown in Figure 1) is controlled by a component in the design of the system. For the two-port directional microphone, it is the damping screen in one of the microphone ports that forms part of an acoustic filter to set the pattern. In the dual-microphone system, the pattern is set by the characteristics of an electrical filter.

The functional similarity of the two methods means they are also alike in

many aspects of their performance. For example, the frequency response curves of Figure 4 apply to both designs. The comments on internal noise and on wind noise later in this lecture also apply equally to both designs. One way in which the systems differ is that the pattern of the two-port directional microphone is set by the microphone manufacturer through the selection of the damping resistor at the time of manufacture. In the dual-microphone system, it is possible to change the characteristics of the electrical filter under control of the DSP to obtain different directional patterns in different situations.

Thus, a hearing aid can provide a number of patterns that can be selected under manual or program control, or it can provide an adaptive directional system that continually modifies the pattern with the objective of maximizing the SNR in a changing environment.

MICROPHONE MATCHING:

One important consideration in the design of a hearing aid with two omnidirectional microphones is unique to that design. When there are two microphones, their inherent sensitivities must be well matched for the patterns to have their intended shapes. Significant microphone

mismatch can degrade the patterns in unintended ways. Hearing aid manufacturers can do several things to ensure that the microphones are well matched. If they are perfectly matched, the signals from the two microphones should have the same overall level and frequency response.

DISADVANTAGES OF DIRECTIONALITY:

To this point, we have not mentioned any disadvantages of directional patterns in hearing aids. However, there are two undesirable features of directional patterns, which in certain circumstances make it best to switch the hearing aid to an omnidirectional mode.

- Internal noise

One disadvantage is internal noise from the hearing aid. In a quiet environment, internal noise may be noticeable. Several studies have shown that the low level of internal noise that is inevitably generated in the microphone and amplifier circuits is not normally audible to the

hearing-impaired user. However, these studies have assumed that an omnidirectional microphone is used in the aid. When a directional microphone or a pair of matched microphones is used to form a directional pattern, the sensitivity falls (see Figure 4). The sensitivity affects

the level of the sensed signal, but the internal noise level is approximately the same as that of an omnidirectional microphone.

Consequently, the signal-to-noise ratio (SNR) at low frequencies is significantly lower in a directional pattern. A hearing aid that uses low-frequency amplification to enhance the audibility of low-frequency signals will naturally amplify the low-frequency internal noise as well. Figure 4 indicates that perhaps 20 dB of additional amplification is needed in a directional pattern to achieve the same level of audibility as in an omnidirectional microphone. This amount of amplification is quite likely to raise the level of internal noise above the environmental ambient noise in a quiet environment.

Thus, the omnidirectional pattern is more appropriate in quiet. In a noisy environment where the benefits of the directionality are needed, the internal noise of even the directional pattern will be well below the ambient level. In hearing aids that allow the user to control the directional mode manually, the user must be counseled that the directional mode should generally not be used in quiet situations.

Some hearing aids provide for automatic switching from the

omnidirectional to the directional mode in backgrounds where the

noise level exceeds a manufacturer determined threshold.

Wind noise:

The presence of wind noise can create another situation in which directional microphone patterns seem noisier than omnidirectional patterns. In this case, the noise is not generated by the microphone, but originates in the turbulent air flow as wind moves past the head and shoulders. We all hear wind noise in even a moderate breeze, and it can be quite loud in many situations. The important feature is that the noise is mostly generated within 10 cm of the ear.

This wind noise presents two difficulties for hearing aid users. First, the microphone is located very close to the source of the noise, while for normal hearers the eardrum is farther from the source and shielded somewhat from it at the end of the ear canal. This affects the signal level of all microphones, whether or not they are directional.

Secondly, directional patterns are relatively more sensitive to sounds from the near field than to sounds from farther away. This is why wind noise is received at a much higher level in directional microphone patterns than in omnidirectional patterns. Also, note that the wind noise effect is the same in a dual-microphone directional implementation as in a two-port directional microphone. For the directional hearing aid user, this means that the aids should be switched to the non-directional mode in wind noise.

Amplifiers:

The basic function of an amplifier is simply to make a small electrical signal into a larger electrical signal. The job of the amplifier is to take power from the battery and transfer it to the amplifier output in a manner controlled by the input signal.

Amplifiers cannot produce signals larger in voltage than some specified maximum. This maximum is usually equal to, or related to, the battery voltage. If the biggest signal in the amplifier is near this maximum, and either the input signal level or the gain of the amplifiers increased, then the amplifier will clip (remove) the peaks of the signal.

The next figure shows the output waveform that results from peak clipping when the input signal is a sine wave.  The thin line shows what the output signal would be if no peak clipping occurred. Because the output is no longer a sine wave, it contains components at frequencies not in the input signal. These additional components are called distortion products.

Tone Controls:

Tone controls in hearing aids have the same function as tone controls in stereos. They cause the gain of the amplifier to vary with frequency.

Filters:

An audio filter is a type of filter used for processing sound signals. Many types of filters exist for applications including graphic equalizers, synthesizers, sound effects, CD players and virtual reality systems.

In its simplest form, an audio filter is typically designed to pass some frequency regions through unattenuated while significantly attenuating others.

The frequency at which attenuation begins is called the cutoff frequency.

3 main types of filters:

- 1. High Pass Filter- is a filter that passes high frequencies well, but attenuates (reduces the amplitude of) frequencies lower than the cutoff frequency. The actual amount of attenuation for each frequency varies from filter to filter.

- 2. A low-pass filter- is a filter that passes low-frequency signals but attenuates signals with frequencies higher than the cutoff frequency.

- 3. A band-pass filter is a device that passes frequencies within a certain range and rejects (attenuates) frequencies outside that range.

- A high-pass filter is the opposite of a low-pass filter, and a band-pass filter is a combination of a high-pass and a low-pass.

Receivers:

The receiver converts the amplified and modified electrical signal into an acoustic output signal.

Telecoils:

A telecoil is a small coil of wire that produces a voltage when an alternating magnetic field flows through it.

Batteries:

The battery provides the increased signal power that the hearing aid delivers to the aid wearer.  The batteries most commonly used for hearing aids use Zinc and Oxygen as their negative and positive electrodes, so the batteries are known as Zinc-air batteries.