Lecture 6

- What types of hearing aid technology are used today?

Hearing aids are distinguished by their technology or circuitry. In the early days, hearing aid technology involved vacuum tubes and large heavy batteries. Today, there are microchips, computerization, and digitized sound processing, used in hearing aid design.

Hearing aids are built from analog or digital circuits. Each technology processes sound differently. Until recently, all hearing aids were analog. Digital hearing aids are the newest kind of hearing aid and are superior to analog.

- Conventional analog hearing aids are designed with a particular frequency response based on the audiogram. The audiologist tells the manufacturer what settings to install. Although there are some adjustments, the aid essentially amplifies all sounds (speech and noise) in the same way. This technology is the least expensive.

- Analog programmable hearing aids have a microchip which allows the aid to have settings programmed for different listening environments such as quiet conversation in the home, noisy situations like a restaurant, or large areas like a theater. The audiologist uses a computer to program the hearing aid for different listening situations depending on the individual hearing loss profile, speech understanding, and range of tolerance for louder sounds.

Some aids can store several programs. As the listening environment changes, the hearing aid settings can be changed by pushing a button on the hearing aid or by using a remote control to switch channels. These aids are more expensive than conventional analog hearing aids, but generally have a longer life span and may provide better hearing in different listening situations.

- Digital programmable hearing aids have all the features of analog programmable aids but use "digitized sound processing" to convert sound waves into digital signals. A computer chip in the aid analyzes the signals of the environment to determine if the sound is noise or speech and then makes modifications to provide a clear, amplified distortion-free signal. Digital hearing aids are usually self-adjusting. The digital processing allows for more flexibility in programming the aid so that the sound it transmits matches the specific pattern of hearing loss. This digital technology is the most expensive, but it allows for improvement in programmability, greater precision in fitting, management of loudness discomfort, control of acoustic feedback (whistling sounds), and noise reduction.

Analog vs. Digital:

All hearing aids, whether analog or digital, are designed to increase the loudness of sounds reaching the ear drum so that the hearing-impaired person can better understand speech. To accomplish this, hearing aids require three basic components:

1. A microphone to gather acoustic energy (sound waves in the air) and convert it to electrical energy.
2. An amplifier to increase the strength of the electrical energy.
3. A receiver, which is like a miniature speaker,  that converts the electrical energy back into acoustic energy (sound waves).

What makes digital hearing aids different from analog hearing aids is what happens in between.

Analog Hearing Aids:

Analog hearing aids use a continuously varying electrical signal to produce sound, just like a microphone and loudspeaker. Analog hearing aids have a microphone that picks up sound and converts the sound into small electrical signals. These signals vary according to the pattern of the sound. The signals are then amplified (made louder) by transistors and fed to the ear phone on the hearing aid which is next to the ear drum.

Most of the better analog hearing aids compress the sound using 'automatic gain control" (AGC). This amplifies quiet sounds until they are loud enough to be heard, but gives less amplification to sounds that are already loud, so as to protect against uncomfortably loud sound levels. Analog hearing aids don’t have all the features that come with advanced digital aids, but they are the least expensive hearing aids available.

Digital Hearing Aids:

Digital aids work in a different way than analog. Digital hearing aids take the signal from the microphone and convert it into "bits" of data - ("0s" and "1s") - numbers that can be manipulated by a tiny computer in the hearing aid. This makes it possible to tailor and process sounds very precisely, in ways that are impossible with analog aids. The bits representing the sound are analyzed and manipulated by algorithms (a set of instructions) to perform precise, complex actions, and are then converted back into electricity, which is finally changed back into sound that goes into the ear. This process happens very rapidly: there are several million calculations occurring in the hearing aid per second. The numbers can be manipulated in almost any way imaginable, and this is what gives the digital hearing aid its big advantage. The binary numbers can perform numerous complex calculations that create very precise, very flexible hearing aids.

## The Benefits of Digital Technology

Digital hearing instruments are truly a breakthrough in hearing technology. The ability to manipulate the sound digitally with a microprocessor (computer) creates hearing aids that offer more features and greater flexibility than analog instruments. Digital manipulation also allows each hearing instrument to be more accurately programmed for each individual hearing loss.

### Cleaner Sound Quality

Digital Signal Processing (DSP) also produces a sharper, cleaner sound quality than analog processing (think about the sound quality of CDs versus audio cassettes).

### Smart Hearing Aids - Background Noise Reduction

Sometimes called "smart" hearing aids, digital instruments have the ability to process sound intelligently. They analyze the incoming sound frequencies and volume. Based on these measurements the digital hearing instrument automatically adjusts to help prevent unpleasant microphone noise or ongoing background noise, such as traffic, while enhancing speech.

### Channels

Channels give more flexibility to program the hearing instruments to the patient’s specific listening needs. The more channels a hearing instrument has, the better the audiologist can fine-tune each sound frequency to best process sound for the patient.

### Feedback Management

Feedback is a high-pitched whistling noise that hearing aid wearers sometimes experience. The primary purpose of feedback management is to reduce the gain (how much sound is amplified) in the frequency ranges in which the feedback is occurring. The result is a reduction of feedback while maintaining clear speech at an adequate volume level.

### Programs/Memories

Multiple programs or memories allow the audiologist to customize a program or memory for different listening situations. You can actually flip a switch to a program designed specifically for quiet or noisy situations, or even for talking on the phone. It can change automatically.

### Volume Control

Volume control helps you adapt to more listening environments. You will need different volumes of sound in a crowded theater than at home sitting across the table from a friend. Some digital hearing devices have an automatic volume control and some have a manual volume control switch.

### Expansion

Expansion, which is the opposite of compression, reduces gain for very soft inputs. In other words, it reduces the microphone noise or background noises (such as the hum of a refrigerator) that can bother hearing aids wearers. Expansion can't be offered in analog hearing aids because the required circuitry is too complex.

### Directional Microphones

Most hearing aid microphones are omnidirectional, which means they pick up sounds from all directions. Directional microphones, on the other hand, are more sensitive to sounds coming from the front, and reduce the sounds coming from behind (often background noise). The result in an increased signal-to-noise ratio that improves listening in noise. While directional microphones can be found on both analog and digital hearing aids, the digital advantage is that the pattern of sensitivity can change depending on the location of the noise, making digital directional technology even more effective.