## Slide 1: Interaction of radiation & matter
Electromagnetic radiation in different regions of spectrum can be used for qualitative and quantitative information
Different types of chemical information

## Slide 2: Energy transfer from photon to molecule or atom
At room temperature most molecules are at lowest electronic & vibrational state
IR radiation can excite vibrational levels that then lose energy quickly in collisions with surroundings

## Slide 3: UV Visible Spectrometry
absorption - specific energy
emission - excited molecule emits
fluorescence
phosphorescence

## Slide 4: What happens to molecule after excitation
collisions deactivate vibrational levels (heat)
emission of photon (fluorescence)
intersystem crossover (phosphorescence)

## Slide 5: General optical spectrometer
Wavelength separation
Photodetectors
Light source - hot objects produce “black body radiation

## Slide 6: Black body radiation
Tungsten lamp, Globar, Nernst glower
Intensity and peak emission wavelength are a function of Temperature
As T increases the total intensity increases and there is shift to higher energies (toward visible and UV)

## Slide 7: UV sources
Arc discharge lamps with electrical discharge maintained in appropriate gases
Low pressure hydrogen and deuterium lamps
Lasers - narrow spectral widths, very high intensity, spatial beam, time resolution, problem with range of wavelengths
Discrete spectroscopic- metal vapor & hollow cathode lamps

## Slide 8: Why separate wavelengths?
Each compound absorbs different colors (energies) with different probabilities (absorbtivity)
Selectivity
Quantitative adherence to Beer’s Law A = abc
Improves sensitivity

## Slide 9: Why are UV-Vis bands broad?
Electronic energy states give band with no vibrational structure
Solvent interactions (microenvironments) averaged
Low temperature gas phase molecules give structure if instrumental resolution is adequate

## Slide 10: Wavelength Dispersion
prisms (nonlinear, range depends on refractive index)
gratings (linear, Bragg’s Law, depends on spacing of scratches, overlapping orders interfere)
interference filters (inexpensive)

## Slide 11: Monochromator
Entrance slit - provides narrow optical image
Collimator - makes light hit dispersive element at same angle
Dispersing element - directional
Focusing element - image on slit
Exit slit - isolates desired color to exit

## Slide 12: Resolution
The ability to distinguish different wavelengths of light - R=
Linear dispersion - range of wavelengths spread over unit distance at exit slit
Spectral bandwidth - range of wavelengths included in output of exit slit (FWHM)
Resolution depends on how widely light is dispersed & how narrow a slice chosen

## Slide 13: Filters - inexpensive alternative
Adsorption type - glass with dyes to adsorb chosen colors
Interference filters - multiple reflections between 2 parallel reflective surfaces - only certain wavelengths have positive interferences - temperature effects spacing between surfaces

## Slide 14: Wavelength dependence in spectrometer
Source
Monochromator
Detector
Sample - We hope so!

## Slide 15: Photodetectors - photoelectric effect E(e)=h
For sensitive detector we need a small work function - alkali metals are best
Phototube - electrons attracted to anode giving a current flow proportional to light intensity
Photomultiplier - amplification to improve sensitivity (10 million)

## Slide 16: Spectral sensitivity is a function of photocathode material
Ag-O-Cs mixture gives broader range but less efficiency
Na2KSb(trace of Cs)has better response over narrow range
Max. response is 10% of one per photon (quantum efficiency)
300nm 500 700 900
Na2KSb
AgOCs

## Slide 17: Photomultiplier - dynodes of CuO.BeO.Cs or GaP.Cs

## Slide 18
Cooled Photomultiplier
Tube

## Slide 19
Dynode array

## Slide 20: Photodiodes - semiconductor that conducts in one direction only when light is present
Rugged and small
Photodiode arrays - allows observation of a number of different locations (wavelengths) simultaneously
Somewhat less sensitive than PMT

## Slide 22: T=I/Io A= - log T = -log (I/Io) Calibration curve

## Slide 24: Deviations from Beer’s Law
High concentrations (0.01M) distort each molecules electronic structure & spectra
Chemical equilibrium
Stray light
Polychromatic light
Interferences

## Slide 25: Interpretation - quantitative
Broad adsorption bands - considerable overlap
Specral dependence upon solvents
Resolving mixtures as linear combinations - need to measure as many wavelengths as components
Beer’s Law .html

## Slide 26: Resolving mixtures
Measure at different wavelengths and solve mathematically
Use standard additions (measure A and then add known amounts of standard)
Chemical methods to separate or shift spectrum
Use time resolution (fluorescence and phosphorescence)

## Slide 27: Improving resolution in mixtures
Instrumental (resolution)
Mathematical (derivatives)
Use second parameter (fluorescence)
Use third parameter (time for phosphorescence)
Chemical separations (chromatography)

## Slide 28: Fluorescence
Emission at lower energy than absorption
Greater selectivity but fluorescent yields vary for different molecules
Detection at right angles to excitation
S/N is improved so sensitivity is better
Fluorescent tags

## Slide 29
Spectrofluorometer
Light source
Monochromator to select excitation
Sample compartment
Monochromator to
select fluorescence

## Slide 30: Photoacoustic spectroscopy
Edison’s observations
If light is pulsed then as gas is excited it can expand (sound)

## Slide 32: Principles of IR
Absorption of energy at various frequencies is detected by IR
plots the amount of radiation transmitted through the sample as a function of frequency
compounds have “fingerprint” region of identity
Notes: <number>
<footer>
<header>
<date/time>

## Slide 33: Infrared Spectrometry
Is especially useful for qualitative analysis
functional groups
other structural features
establishing purity
monitoring rates
measuring concentrations
theoretical studies

## Slide 34: How does it work?
Continuous beam of radiation
Frequencies display different absorbances
Beam comes to focus at entrance slit
molecule absorbs radiation of the energy to excite it to the vibrational state
Notes: <number>
<footer>
<header>
<date/time>

## Slide 35: How Does It Work?
Monochromator disperses radiation into spectrum
one frequency appears at exit slit
radiation passed to detector
detector converts energy to signal
signal amplified and recorded
Notes: <number>
<footer>
<header>
<date/time>

## Slide 36: Instrumentation II
Optical-null double-beam instruments
Radiation is directed through both cells by mirrors
sample beam and reference beam
chopper
diffraction grating

## Slide 37: Double beam/ null detection

## Slide 38: Instrumentation III
Exit slit
detector
servo motor
Resulting spectrum is a plot of the intensity of the transmitted radiation versus the wavelength

## Slide 39: Detection of IR radiation
Insufficient energy to excite electrons & hence photodetectors won’t work
Sense heat - not very sensitive and must be protected from sources of heat
Thermocouple - dissimilar metals characterized by voltage across gap proportional to temperature

## Slide 40: IR detectors
Golay detector - gas expanded by heat causes flexible mirror to move - measure photocurrent of visible light source
Detector
IR beam
Vis
source
GAS
Flexible mirror

## Slide 41: Carbon analyzer - simple IR
Sample flushed of carbon dioxide (inorganic)
Organic carbon oxidized by persulfate & UV
Carbon dioxide measured in gas cell (water interferences)

## Slide 42: NDIR detector - no monochromator
Chopper
SAMP
REF
Detector cell
Filter
CO2
CO2
Beam trimmer
Press. sens. det.

## Slide 43: Limitations
Mechanical coupling
Slow scanning / detectors slow

## Slide 44: Limitations of Dispersive IR
Mechanically complex
Sensitivity limited
Requires external calibration
Tracking errors limit resolution (scanning fast broadens peak, decreases absorbance, shifts peak

## Slide 45: Problems with IR
c no quantitative
H limited resolution
D not reproducible
A limited dynamic range
I limited sensitivity
E long analysis time
B functional groups

## Slide 46: Limitations
Most equipment can measure one wavelength at a time
Potentially time-consuming
A solution?

## Slide 47: Fourier-Transform Infrared Spectroscopy (FTIR)
A Solution!

## Slide 48: FTIR
Analyze all wavelengths simultaneously
signal decoded to generate complete spectrum
can be done quickly
better resolution
more resolution
However, . . .

## Slide 49: FTIR
A solution, yet an expensive one!
FTIR uses sophisticated machinery more complex than generic GCIR

## Slide 50: Fourier Transform IR
Mechanically simple
Fast, sensitive, accurate
Internal calibration
No tracking errors or stray light

## Slide 51: IR Spectroscopy - qualitative
Double beam required to correct for blank at each wavelength
Scan time (sensitivity) Vs resolution
Michelson interferometer & FTIR

## Slide 52: Advantages of FTIR
Multiplex--speed, sensitivity (Felgett)
Throughput--greater energy, S/N (Jacquinot)
Laser reference--accurate wavelength, reproducible (Connes)
No stray light--quantitative accuracy
No tracking errors--wavelength and photometric accuracy

## Slide 53: New FTIR Applications
Quality control--speed, accuracy
Micro, trace analysis--nanogram levels, small samples
Kinetic studies--milliseconds
Internal reflection
Telescopic

## Slide 54: Attenuated Internal Reflection
Surface analysis
Limited by 75% energy loss

## Slide 55: New FTIR Applications
Quality control--speed, accuracy
Micro, trace analysis--nanogram levels, small samples
Kinetic studies--milliseconds
Internal reflection
Telescopic